# U-Pb geochronologic analyses of zircon (Element2 HR ICPMS)

Zircon crystals are extracted from samples by traditional methods of crushing and grinding, followed by separation with a Wilfley table, heavy liquids, and a Frantz magnetic separator. Samples are processed such that all zircons are retained in the final heavy mineral fraction. For detrital analyses, a large split of grains (generally thousands of grains) is incorporated into a 1” epoxy mount together with fragments or loose grains of Sri Lanka, FC-1, and R33 zircon crystals that are used as primary standards. For igneous samples, ~50 high-quality grains are selected and mounted with standards, generally with four samples per mount. The mounts are sanded down to a depth of ~20 microns, polished, imaged, and cleaned prior to isotopic analysis.

U-Pb geochronology of zircons is conducted by laser ablation inductively coupled plasma mass spectrometry (LA-ICPMS) at the Arizona LaserChron Center (Gehrels et al., 2006, 2008; Gehrels and Pecha, 2014). The analyses involve ablation of zircon with a Photon Machines Analyte G2 excimer laser equipped with HelEx ablation cell using a spot diameter of 20 microns. The ablated material is carried in helium into the plasma source of an Element2 HR ICPMS, which sequences rapidly through U, Th, and Pb isotopes. Signal intensities are measured with an SEM that operates in pulse counting mode for signals less than 50K cps, in both pulse-counting and analog mode for signals between 50K and 5M cps, and in analog mode above 5M cps. The calibration between pulse-counting and analog signals is determined line-by-line for signals between 50K and 5M cps, and is applied to >5M cps signals. Four intensities are determined and averaged for each isotope, with dwell times of 0.0052 sec for 202, 0.0075 sec for 204, 0.0202 sec for 206, 0.0284 sec for 207, 0.0026 sec for 208, 0.0026 sec for 232, and 0.0104 sec for 238.

With the laser set an energy density of ~5 J/cm2, a repetition rate of 8 hz, and an ablation time of 10 seconds, ablation pits are ~12 microns in depth. Sensitivity with these settings is approximately ~5,000 cps/ppm. Each analysis consists of 5 sec on peaks with the laser off (for backgrounds), 10 sec with the laser firing (for peak intensities), and a 20 second delay to purge the previous sample and save files.

Prior to analysis, grains are imaged to provide a guide for locating analysis pits in optimal locations, and to assist in interpreting results. Images are made with a Hitachi 3400N SEM and a Gatan CL2 detector system (www.geoarizonasem.org). In general, BSE images are made for detrital mounts and CL images are made for igneous mounts.

Following analysis, data reduction is performed with an in-house Python decoding routine and an Excel spreadsheet (E2agecalc) that:

1. Decodes .dat files from the Thermo software such individual intensities for measurement are available (routine written by John Hartman, University of Arizona)

2. Imports intensities and a sample name for each analysis

3. Calculates average intensities for each isotope (based on the sum of all counts while the laser is firing)

4. Subtracts 204Hg from the 204 signal to yield 204Pb intensity (using natural 202Hg/204Hg of 4.3). This Hg correction is not significant for most analyses because our Hg backgrounds are low (generally ~150 cps at mass 204).

5. Performs a common Pb correction based on the measured 206Pb/204Pb and the assumed composition of common Pb based on Stacey and Kramers (1975)

6. Calculates measured 206/238, 206/207, and 208/232 ratios

7. Compares measured and known ratios for the three standards to determine fractionation factors for 206/238, 206/207, and 208/232. These correction factors are generally <5% for 206/238, <2% for 206/207, and <20% for 208/232.

8. Determines an overdispersion factor if the standard analyses show greater dispersion than expected from measurement uncertainties

9. Uses a sliding-window average to apply fractionation factors to unknowns (generally averaging 8 standard analyses)

10. Calculates fractionation-corrected 206/238, 206/207, and 208/232 ratios and ages for unknowns

11. Propagates measurement uncertainties for 206/238 and 208/232 that are based on the scatter about a regression of measured values. Uncertainties for 206/207 and 206/204 are based on the standard deviation of measured values since these ratios generally do not change during an analysis. The sum of this uncertainty and any overdispersion factor is reported as the internal (or measurement) uncertainty for each analysis. These uncertainties are reported at the 1-sigma level.

12. Calculates the down-hole slope of 206/238 to highlight analyses in which 206/238 is compromised due to heterogeneity in age (e.g., crossing an age boundary) or intersection of a fracture or inclusion.

13. Calculates concentrations of U and Th for unknowns based on the measured intensity and known concentrations of FC-1.

14. Calculates the external (systematic) uncertainties for 206/238, 206/207, and 208/232, which include contributions from (a) the scatter of standard analyses, (b) uncertainties in the ages of the standards, (c) uncertainties in the composition of common Pb, and (4) uncertainties in the decay constants for 235U and 238U.

15. Determines a “Best Age” for each analysis, which is generally the 206/238 age for <900 Ma ages and the 206/207 age for >900 Ma ages.

16. Provides preliminary filters that highlight analyses with >20% discordance, >5% reverse discordance, or >10% internal (measurement) uncertainty.

17. Corrects 206/238U ages for U-Th disequilibrium. This has a significant impact only on very young (~<2 Ma) ages.

18. Calculates the radiation dosage that the analyzed portion of each zircon has experienced, assuming a value of 2.3 for the Th/U of the magma. This is plotted against 206/238 age to help identify Pb loss.

18. Creates a publication-ready datatable with concentrations, isotope ratios, and ages for unknowns.

For detrital analyses, the ages are shown on Pb*/U concordia diagrams and relative age-probability diagrams using the routines in Isoplot (Ludwig, 2008). The age-probability diagrams show each age and its uncertainty (for measurement error only) as a normal distribution, and sum all ages from a sample into a single curve. Composite age probability plots are made from an in-house Excel program (see Analysis Tools for link) that normalizes each curve according to the number of constituent analyses, such that each curve contains the same area, and then stacks the probability curves.

For igneous analyses, the ages are shown on Pb*/U concordia diagrams and weighted mean diagrams using the routines in Isoplot (Ludwig, 2008). The weighted mean diagrams show the weighted mean (weighting according to the square of the internal uncertainties), the uncertainty of the weighted mean, the external (systematic) uncertainty that corresponds to the ages used, the final uncertainty of the age (determined by quadratic addition of the weighted mean and external uncertainties), and the MSWD of the data set.

References:

Gehrels, G.E., Valencia, V., Ruiz, J., 2008, Enhanced precision, accuracy, efficiency, and spatial resolution of U-Pb ages by laser ablation–multicollector–inductively coupled plasma–mass spectrometry: Geochemistry, Geophysics, Geosystems, v. 9, Q03017, doi:10.1029/2007GC001805.

Gehrels, G.E., Valencia, V., Pullen, A., 2006, Detrital zircon geochronology by Laser-Ablation Multicollector ICPMS at the Arizona LaserChron Center, in Loszewski, T., and Huff, W., eds., Geochronology: Emerging Opportunities, Paleontology Society Short Course: Paleontology Society Papers, v. 11, 10 p.

Gehrels, G. and Pecha, M., 2014, Detrital zircon U-Pb geochronology and Hf isotope geochemistry of Paleozoic and Triassic passive margin strata of western North America: Geosphere, v. 10 (1), p. 49-65.

Ludwig, K.R., 2008, Isoplot 3.60. Berkeley Geochronology Center, Special Publication No. 4, 77 p.

Stacey, J.S., and Kramers, J.D., 1975, Approximation of terrestrial lead isotope evolution by a two-stage model: Earth and Planetary Science Letters, v. 26, p. 207-221.

****************************************************************************

Notes to be inserted below data table:

- 1. Analyses with >10% uncertainty (1-sigma) in 206Pb/238U age are not included.

- Analyses with >10% uncertainty (1-sigma) in 206Pb/207Pb age are not included, unless 206Pb/238U age is <400 Ma.

- Best age is determined from 206Pb/238U age for analyses with 206Pb/238U age <900 Ma and from 206Pb/207Pb age for analyses with 206Pb/238Uage >900 Ma.

- Concordance is based on 206Pb/238U age / 206Pb/207Pb age.  Value is not reported for 206Pb/238U ages <400 Ma because of large uncertainty in 206Pb/207Pb age.

- Analyses with 206Pb/238U age >400 Ma and with >20% discordance (<80% concordance) are not included.

- Analyses with 206Pb/238U age >400 Ma and with >5% reverse discordance (<105% concordance) are not included.

- All uncertainties are reported at the 1-sigma level, and include only measurement errors.

- Systematic errors are as follows (at 2-sigma level): [sample 1: xxx% (206Pb/238U) & xxx% (206Pb/207Pb)] These values are reported on cells U1 and W1 of E2agecalc.

- Analyses conducted by LA-ICPMS, as described by Gehrels et al. (2008) and Gehrels and Pecha (2014).

- U concentration and U/Th are calibrated relative to FC-1 zircon standard and are accurate to ~20%.

- Common Pb correction is from measured 204Pb with common Pb composition interpreted from Stacey and Kramers (1975).

- Common Pb composition assigned uncertainties of 1.5 for 206Pb/204Pb, 0.3 for 207Pb/204Pb, and 2.0 for 208Pb/204Pb.

- U/Pb and 206Pb/207Pb fractionation is calibrated relative to fragments of large Sri Lanka zircons and individual crystals of FC-1, and R33.

- U decay constants and composition as follows: 238U = 9.8485 x 10-10, 235U = 1.55125 x 10-10, 238U/235U = 137.82.

- U-Th disequilibrium correction is applied to 206/238 ages assuming a value of 2.3 for the magma.

- Weighted mean and concordia plots determined with Isoplot (Ludwig, 2008).

References:

Gehrels, G.E., Valencia, V., Ruiz, J., 2008, Enhanced precision, accuracy, efficiency, and spatial resolution of U-Pb ages by laser ablation-multicollector-inductively coupled plasma-mass spectrometry: Geochemistry, Geophysics, Geosystems, v. 9, Q03017, doi:10.1029/2007GC001805.

Gehrels, G. and Pecha, M., 2014, Detrital zircon U-Pb geochronology and Hf isotope geochemistry of Paleozoic and Triassic passive margin strata of western North America: Geosphere, v. 10 (1), p. 49-65.

Ludwig, K., 2008, Isoplot 3.6: Berkeley Geochrono9logy Center Special Publication 4, 77 p.

Stacey, J.S., and Kramers, J.D., 1975, Approximation of terrestrial lead isotope evolution by a two stage model: Earth and Planetary Science Letters, v. 26, p. 207-221.